EPA TO-17 Volatile Organic Compounds

Method TO-17  is used to analyze samples for volatile organic compounds collected on multi-bed sorbent tubes, which are thermally desorbed and cryo-focused on the capillary column and then analyzed by GC/MS.  The range of compounds analyzed by the method depends on the selection of the sorbent cartridge. EAS follows the method recommendation that the calibration and QC criteria for Method TO-17 follow the TO-15 method.

The modifications done by EAS include the target list.  The method recommends using the Method TO-15 QC criteria.  EAS uses the modified TO-15 QC criteria listed in Table 13.9a.

#### Table 13.9a

#### TO-17 Sorbent Cartridge Selection Guide

| Tube Name | Compounds | Packing | Desorption Temperature |
| --- | --- | --- | --- |
| Tenax TA | BTEX Diesel Range Organic | Tenax TA | 300C |
| Carbotrap 300 | General VOC | Carbopak C Carbopak B Carboseive SIII | 325C |
| VOC | General VOC | Tenax TA Carboxen 1000 Carboseive SIII | 325C |

TO-17 tubes can also be sampled passively using special adapters.  The tubes are desorbed and analyzed in the same manner as the normal TO-17, and the TO-17 QC criteria is used.

#### Table 13.9b

#### TO-17 Recommended Sampling Times

| Final Volume | Flow Rate | Time |
| --- | --- | --- |
| 500 ml | 100 ml/min | 5 min |
| 480 ml | 1 ml/min | 8 hours |
| 720 ml | 0.5 ml/min | 24 hours |

#### Table 13.9c

#### TO-17 Summary of QC Criteria for Method TO-17

| Parameter | EAS TO-15 Modified | TO-17 Method |
| --- | --- | --- |
| BFB Tune | Daily (24 hour) 12 hours if Required | Daily (24 hour) |
| Tuning Criteria with BFB | TO-15 Tune Criteria | TO-15 Tune Criteria |
| Initial Calibration | Five points minimum See Table 13.9b 90%compounds meet criteria | Same as TO-15. Either liquid standards or gas phase standards can be used. |
| Calibration Check Sample (CCS) | After Initial Calibration Same Percent D as Initial Calibration | Same as TO-15 |
| Continuing Calibration Verification (CCV) | Daily (24 hours) See Table 13.9b 90%compounds meet criteria |  |
| Internal Standard (IS) | Pentafluorobenzene 1,4-Difluorobenzene RT < 0.5 min daily std. Response 60% to 140% 20 ppbv | Introduce gas phase internal standard onto sorbent tube (optional) |
| Surrogate | Toluene-d8 70-130% recovery |  |
| Tube Blank | <RL for Tubes from the same media. From client media if supplied | Once tubes are analyzed they can be considered clean and can be reused Artifact peaks should be identified in final report. |
| Laboratory Control Spike | 1 per Daily Batch 70-130% for LCS list See Table 13.7b | Same as TO-15 |
| Duplicate Lab Control Dup Duplicate Pairs | Duplicate Pair with each 20 samples <30% for LCS spike list See Table 13.7b | The precision is tested by using six standard tubes and repeated every 10 series of samples |

The laboratory control spike (LCS) compounds are indicated in bold in Table 13.9d. The MDL for ppbV is based on a sample volume of 5L. This is the maximum volume that should be used.

Table 13.9d

TO-17 Method TO-17 Compound List and QC Criteria

The TO-17 method can be run with any of the 100 VOC compounds listed below, but is generally done using the TO-17 Short List or the TO-17 List shown in Table 13.9e or 13.9f..

|  | ICAL/CCV | LCS | Precision |
| --- | --- | --- | --- |
| Component | %D | %R | %D |
| Freon 12 | <30% | 70-130 | <30% |
| Chloromethane | <30% | 70-130 | <30% |
| Freon 114 | <30% | 70-130 | <30% |
| Vinyl chloride | <30% | 70-130 | <30% |
| Bromomethane | <30% | 70-130 | <30% |
| Chloroethane | <30% | 70-130 | <30% |
| Trichlorofluoromethane | <30% | 70-130 | <30% |
| 1,1-Dichloroethene | <30% | 70-130 | <30% |
| Dichloromethane | <30% | 70-130 | <30% |
| Freon 113 | <30% | 70-130 | <30% |
| 1,1-Dichloroethane | <30% | 70-130 | <30% |
| c-1,2-Dichloroethene | <30% | 70-130 | <30% |
| Chloroform | <30% | 70-130 | <30% |
| 1,2-Dichloroethane | <30% | 70-130 | <30% |
| 1,1,1-Trichloroethane | <30% | 70-130 | <30% |
| Benzene | <30% | 70-130 | <30% |
| Carbon Tetrachloride | <30% | 70-130 | <30% |
| 1,2-Dichloropropane | <30% | 70-130 | <30% |
| Trichloroethene | <30% | 70-130 | <30% |
| c-1,3-Dichloropropene | <30% | 70-130 | <30% |
| t-1,3-Dichloropropene | <30% | 70-130 | <30% |
| 1,1,2-Trichloroethane | <30% | 70-130 | <30% |
| Toluene | <30% | 70-130 | <30% |
| 1,2-Dibromoethane | <30% | 70-130 | <30% |
| Tetrachloroethene | <30% | 70-130 | <30% |
| Chlorobenzene | <30% | 70-130 | <30% |
| Ethylbenzene | <30% | 70-130 | <30% |
| M,p-Xylenes | <30% | 70-130 | <30% |
| Styrene | <30% | 70-130 | <30% |
| o-Xylene | <30% | 70-130 | <30% |
| 1,1,2,2-Tetrachloroethane | <30% | 70-130 | <30% |
| 1,3,5-Trimethylbenzene | <30% | 70-130 | <30% |
| 1,2,4-Trimethylbenzene | <30% | 70-130 | <30% |
| 1,3-Dichlorobenzene | <30% | 70-130 | <30% |
| 1,4-Dichlorobenzene | <30% | 70-130 | <30% |
| 1,2-Dichlorobenzene | <30% | 70-130 | <30% |
| 1,2,4-Trichlorobenzene | <30% | 70-130 | <30% |
| Hexachlorobutadiene | <50% | 50-150 | <50% |
| TO-15 Compounds |  |  |  |
| 1,3-Butadiene | <40% | 60-140 | <40% |
| 2-Butanone | <40% | 60-140 | <40% |
| Acetone | <40% | 60-140 | <40% |
| Carbon Disulfide | <40% | 60-140 | <40% |
| Bromoform | <40% | 60-140 | <40% |
| 4-Methyl-2-pentanone | <40% | 60-140 | <40% |
| 2-Hexanone | <40% | 60-140 | <40% |
| Bromodichloromethane | <40% | 60-140 | <40% |
| Dibromochloromethane | <40% | 60-140 | <40% |
| Vinyl acetate | <40% | 60-140 | <40% |
| t-1,2-Dichloroethene | <40% | 60-140 | <40% |
| Benzylchloride | <50% | 50-150 | <50% |
| 4-Ethyltoluene | <40% | 60-140 | <40% |
| Methyl t-butyl ether | <40% | 60-140 | <40% |
| Cyclohexane | <40% | 60-140 | <40% |
| 1,4-Dioxane | <40% | 60-140 | <40% |
| Tetrahydrofuran | <40% | 60-140 | <40% |
| Hexane | <40% | 60-140 | <40% |
| Heptane | <40% | 60-140 | <40% |
| 2,2,4-Trimethylpentane | <40% | 60-140 | <40% |
| 3-Chloroprene | <40% | 60-140 | <40% |
| Ethyl-Acetate | <40% | 60-140 | <40% |
| 2-Propanol | <40% | 60-140 | <40% |
| TO-15 8260 Compounds |  |  |  |
| n-Proplylbenzene | <40% | 60-140 | <40% |
| Isopropylbenzene | <40% | 60-140 | <40% |
| 2,2-Dichloropropane | <40% | 60-140 | <40% |
| 1,1,1,2-Tetrachloroethane | <40% | 60-140 | <40% |
| Bromochloromethane | <40% | 60-140 | <40% |
| Octane | <40% | 60-140 | <40% |
| Nonane | <40% | 60-140 | <40% |
| Decane | <40% | 60-140 | <40% |
| 1,1-Dichloropropene | <40% | 60-140 | <40% |
| 1,2,3 Trichloropropane | <40% | 60-140 | <40% |
| 1,3-Dichloropropane | <40% | 60-140 | <40% |
| Dibromomethane | <40% | 60-140 | <40% |
| Methyl methacrylate | <40% | 60-140 | <40% |
| Di-isopropyl ether | <40% | 60-140 | <40% |
| Isobutyl Alcohol | <40% | 60-140 | <40% |
| Ethanol | <40% | 60-140 | <40% |
| n-Butylbenzene | <40% | 60-140 | <40% |
| sec-Butylbenzene | <40% | 60-140 | <40% |
| tert-butylbenzene | <40% | 60-140 | <40% |
| i-Butylbenzene | <40% | 60-140 | <40% |
| p-Isopropyltoluene | <40% | 60-140 | <40% |
| t-Butanol | <40% | 60-140 | <40% |
| 2-Chlorotoluene | <40% | 60-140 | <40% |
| 4-Chlorotoluene | <40% | 60-140 | <40% |
| Methyl Acrylate | <40% | 60-140 | <40% |
| Ethyl tert butyl Ether | <40% | 60-140 | <40% |
| 1,2,3-Trichlorobenzene | <40% | 60-140 | <40% |

Table 13.9e

Method TO-17 Short Compound List and MDL

This is a list of common air pollutants that can be analyzed by TO-17.  The MDL for TO-17 is reported in ug per sample.  The actual MDL depends on the air volume sampled.  The MDL in ppbv and ug/m3 is based on a 1,000 ml air volume.

| ###### CAS | Compound | MDL ug | MDL ppbV | RL ppbv | MDL ug/m3 | RL ug/m3 |
| --- | --- | --- | --- | --- | --- | --- |
| 75-01-4 | Vinyl chloride | 0.3 | 0.5 | 1.0 | 1.3 | 2.6 |
| 75-35-4 | 1,1-Dichloroethene | 0.4 | 0.5 | 1.0 | 2.0 | 4.0 |
| 76-13-1 | Freon 113 | 0.8 | 0.5 | 1.0 | 3.8 | 7.7 |
| 75-09-2 | Methylene Chloride | 0.3 | 0.5 | 1.0 | 1.7 | 3.5 |
| 75-34-3 | 1,1-Dichloroethane | 0.4 | 0.5 | 1.0 | 2.0 | 4.1 |
| 156-59-2 | cis-1,2-Dichloroethene | 0.4 | 0.5 | 1.0 | 2.0 | 4.0 |
| 67-66-3 | Chloroform | 0.5 | 0.5 | 1.0 | 2.4 | 4.9 |
| 71-55-6 | 1,1,1-Trichloroethane | 0.5 | 0.5 | 1.0 | 2.7 | 5.5 |
| 107-06-2 | 1,2-Dichloroethane | 0.5 | 0.5 | 1.0 | 2.0 | 4.1 |
| 71-43-2 | Benzene | 0.3 | 0.5 | 1.0 | 1.6 | 3.2 |
| 56-23-5 | Carbon tetrachloride | 0.6 | 0.5 | 1.0 | 3.2 | 6.3 |
| 78-87-5 | 1,2-Dichloropropane | 0.5 | 0.5 | 1.0 | 2.3 | 4.6 |
| 79-01-6 | Trichloroethene | 0.5 | 0.5 | 1.0 | 2.7 | 5.4 |
| 10061-01-5 | cis-1,3-Dichloropropene | 0.5 | 0.5 | 1.0 | 2.3 | 4.5 |
| 108-88-3 | Toluene | 0.4 | 0.5 | 1.0 | 1.9 | 3.8 |
| 10061-02-6 | trans-1,3-Dichloropropene | 0.5 | 0.5 | 1.0 | 2.3 | 4.5 |
| 79-00-5 | 1,1,2-Trichloroethane | 0.5 | 0.5 | 1.0 | 2.7 | 5.5 |
| 106-93-4 | 1,2-Dibromoethane | 0.8 | 0.5 | 1.0 | 3.9 | 7.7 |
| 127-18-4 | Tetrachloroethene | 0.3 | 0.3 | 0.5 | 1.7 | 3.4 |
| 108-90-7 | Chlorobenzene | 0.5 | 0.5 | 1.0 | 2.3 | 4.6 |
| 100-41-4 | Ethylbenzene | 0.4 | 0.5 | 1.0 | 2.2 | 4.4 |
| 1330-20-7 | m,p-Xylene | 0.4 | 0.5 | 1.0 | 2.2 | 4.4 |
| 95-47-6 | o-Xylene | 0.4 | 0.5 | 1.0 | 2.2 | 4.4 |
| 108-67-8 | 1,3,5-Trimethylbenzene | 0.5 | 0.5 | 1.0 | 2.5 | 4.9 |
| 95-63-6 | 1,2,4-Trimethylbenzene | 0.5 | 0.5 | 1.0 | 2.5 | 4.9 |

Table 13.9f

#### Method TO-17 Compounds with MDL and RL

The TO-17 Compound List with MDL.  For TO-17 the MDL is reported as ug per sample.  The MDL can be calculated by dividing the MDL in ug by the air volume in m3.  The MDL and RL values below are for an air volume of 1,000 ml.  The TO-17 method can be run with any of the 100 VOC compounds listed on the TO-15 Extended list.

| CAS | # Compound | MDL ug | MDL ppbV | # RL ppbv | MDL ug/m3 | RL ug/m3 |
| --- | --- | --- | --- | --- | --- | --- |
| 75-71-8 | Dichlorodifluoromethane | 0.5 | 0.4 | 0.7 | 2 | 4 |
| 74-87-3 | Chloromethane | 0.2 | 0.4 | 0.7 | 1 | 2 |
| 76-14-2 | Freon 114 | 0.7 | 0.4 | 0.7 | 3 | 5 |
| 75-01-4 | Vinyl chloride | 0.3 | 0.4 | 0.7 | 1 | 2 |
| 106-99-0 | 1,3-Butadiene | 0.2 | 0.4 | 0.7 | 1 | 2 |
| 74-83-9 | Bromomethane | 0.4 | 0.4 | 0.7 | 1 | 3 |
| 75-00-3 | Chloroethane | 0.3 | 0.4 | 0.7 | 1 | 2 |
| 75-69-4 | Trichlorofluoromethane | 0.6 | 0.4 | 0.7 | 2 | 4 |
| 67-64-1 | Acetone | 0.2 | 0.4 | 0.7 | 1 | 2 |
| 75-35-4 | 1,1-Dichloroethene | 0.4 | 0.4 | 0.7 | 1 | 3 |
| 76-13-1 | Freon 113 | 0.8 | 0.4 | 0.7 | 3 | 6 |
| 107-05-1 | Allyl chloride | 0.2 | 0.4 | 0.7 | 1 | 2 |
| 75-09-2 | Methylene Chloride | 0.3 | 0.4 | 0.7 | 1 | 3 |
| 75-15-0 | Carbon disulfide | 0.3 | 0.4 | 0.7 | 1 | 2 |
| 156-60-5 | trans-1,2-Dichloroethene | 0.2 | 0.2 | 0.4 | 1 | 1 |
| 1634-04-4 | Methyl tert butyl ether | 0.2 | 0.2 | 0.4 | 1 | 1 |
| 75-34-3 | 1,1-Dichloroethane | 0.4 | 0.4 | 0.7 | 2 | 3 |
| 108-05-4 | Vinyl acetate | 0.4 | 0.4 | 0.7 | 1 | 3 |
| 78-93-3 | 2-Butanone | 0.3 | 0.4 | 0.7 | 1 | 2 |
| 110-54-3 | Hexane | 0.2 | 0.2 | 0.4 | 1 | 1 |
| 141-78-6 | Ethyl acetate | 0.4 | 0.4 | 0.7 | 1 | 3 |
| 109-99-9 | Tetrahydrofuran | 0.3 | 0.4 | 0.7 | 1 | 2 |
| 156-59-2 | cis-1,2-Dichloroethene | 0.4 | 0.4 | 0.7 | 1 | 3 |
| 67-66-3 | Chloroform | 0.5 | 0.4 | 0.7 | 2 | 4 |
| 71-55-6 | 1,1,1-Trichloroethane | 0.5 | 0.4 | 0.7 | 2 | 4 |
| 107-06-2 | 1,2-Dichloroethane | 0.4 | 0.4 | 0.7 | 2 | 3 |
| 563-58-6 | 1,1-Dichloropropene | 0.2 | 0.2 | 0.4 | 1 | 2 |
| 110-82-7 | Cyclohexane | 0.2 | 0.2 | 0.4 | 1 | 1 |
| 71-43-2 | Benzene | 0.3 | 0.4 | 0.7 | 1 | 2 |
| 56-23-5 | Carbon tetrachloride | 0.6 | 0.4 | 0.7 | 2 | 5 |
| 540-84-1 | 2,2,4-Trimethylpentane | 0.2 | 0.2 | 0.4 | 1 | 2 |
| 142-82-5 | n-Heptane | 0.2 | 0.2 | 0.4 | 1 | 2 |
| 78-87-5 | 1,2-Dichloropropane | 0.5 | 0.4 | 0.7 | 2 | 3 |
| 123-91-1 | 1,4 Dioxane | 0.7 | 0.7 | 1.5 | 3 | 5 |
| 79-01-6 | Trichloroethene | 0.5 | 0.4 | 0.7 | 2 | 4 |
| 75-27-4 | Bromodichloromethane | 0.3 | 0.2 | 0.4 | 1 | 3 |
| 108-10-1 | 4-Methyl-2-pentanone | 0.2 | 0.2 | 0.4 | 1 | 2 |
| 10061-01-5 | cis-1,3-Dichloropropene | 0.5 | 0.4 | 0.7 | 2 | 3 |
| 108-88-3 | Toluene | 0.4 | 0.4 | 0.7 | 1 | 3 |
| 10061-02-6 | trans-1,3-Dichloropropene | 0.5 | 0.4 | 0.7 | 2 | 3 |
| 79-00-5 | 1,1,2-Trichloroethane | 0.5 | 0.4 | 0.7 | 2 | 4 |
| 591-78-6 | 2-Hexanone | 0.2 | 0.2 | 0.4 | 1 | 2 |
| 142-28-9 | 1,3-Dichloropropane | 0.2 | 0.2 | 0.4 | 1 | 2 |
| 124-48-1 | Dibromochloromethane | 0.4 | 0.2 | 0.4 | 2 | 3 |
| 106-93-4 | 1,2-Dibromoethane | 0.8 | 0.4 | 0.7 | 3 | 6 |
| 127-18-4 | Tetrachloroethene | 0.3 | 0.2 | 0.4 | 1 | 3 |
| 108-90-7 | Chlorobenzene | 0.5 | 0.4 | 0.7 | 2 | 3 |
| 100-41-4 | Ethylbenzene | 0.4 | 0.4 | 0.7 | 2 | 3 |
| 1330-20-7 | m,p-Xylene | 0.4 | 0.4 | 0.7 | 2 | 3 |
| 100-42-5 | Styrene | 0.4 | 0.4 | 0.7 | 2 | 3 |
| 75-25-2 | Bromoform | 0.3 | 0.1 | 0.2 | 1 | 2 |
| 95-47-6 | o-Xylene | 0.4 | 0.4 | 0.7 | 2 | 3 |
| 79-34-5 | 1,1,2,2-Tetrachloroethane | 0.3 | 0.2 | 0.4 | 1 | 3 |
| 96-18-4 | 1,2,3-Trichloropropane | 0.3 | 0.2 | 0.4 | 1 | 2 |
| 622-96-8 | 4-Ethyltoluene | 0.2 | 0.2 | 0.4 | 1 | 2 |
| 108-67-8 | 1,3,5-Trimethylbenzene | 0.5 | 0.4 | 0.7 | 2 | 4 |
| 95-63-6 | 1,2,4-Trimethylbenzene | 0.5 | 0.4 | 0.7 | 2 | 4 |
| 541-73-1 | 1,3-Dichlorobenzene | 0.3 | 0.2 | 0.4 | 1 | 2 |
| 100-44-7 | Benzyl chloride | 0.3 | 0.2 | 0.4 | 1 | 2 |
| 106-46-7 | 1,4-Dichlorobenzene | 0.3 | 0.2 | 0.4 | 1 | 2 |
| 95-50-1 | 1,2-Dichlorobenzene | 0.3 | 0.2 | 0.4 | 1 | 2 |
| 120-82-1 | 1,2,4-Trichlorobenzene | 0.2 | 0.4 | 0.7 | 3 | 6 |
| 87-68-3 | Hexachlorobutadiene | 0.3 | 0.4 | 0.7 | 4 | 8 |